## Slide 1: A Prototype Device for Facial Electromyography Signal Collection Final Report Date: 05/03/2021 Sponsor: Dr. Rose T. Faghih Project Manager: Dr. Steven Pei
Team 18
Chuong Nguyen (Team Leader) | Toan Vu | Johnathan Nguyen

## Slide 2: PROBLEMS
70% of adults experience at least one traumatic event
90% of those who die by suicide had a diagnosable psychiatric disorder
Low success rate of Psychotherapy treatment (50%)
EMG Device | 05/03/2021
Notes: https://www.scottdmiller.com/the-failure-rate-of-psychotherapy-what-it-is-and-what-we-can-do/

## Slide 3: NEEDS
A device to provide well-measured data to medical professionals
More accurate diagnosis with better treatment courses for patients
EMG Device | 05/03/2021

## Slide 4: PURPOSE
Detect a person’s emotion by collecting and analysing their electromyography (EMG) signals.
Display the user’s emotion on a mobile application.
EMG Device | 05/03/2021

## Slide 5: OVERVIEW DIAGRAM
EMG Device | 05/03/2021

## Slide 6: PATENTS
US 8214214 B2
  Emotion detection by analyzing the user’s prosody (rhythm and intonation of voice)
US 9600715 B2
  Emotion detection by analyzing raw image data of a user’s facial expression
US 8170656 B2
  Detecting the gestures of parts of body by analyzing the EMG signals
EMG Device | 05/03/2021

## Slide 7: DELIVERABLES
Microcontroller Programing
iOS Mobile Application
EMG Device | 05/03/2021
Machine Learning Model

## Slide 8: ENGINEERING SPECIFICATIONS
Bluetooth Connectivity | Bluetooth Low Energy (BLE) Range: 2-3 meters
Sampling Frequency | 1 [KHz]
Sensors | 3 Myoware Muscle Sensors
Power Consumption | 127 [mAh]
Machine Learning | TensorFlow, Recurrent Neural Network, Long-Short Term Memory
Machine Learning Model | 90% Accuracy
Emotion Predictions | 70% Accuracy
EMG Device | 05/03/2021
Notes: Tensorflow is a Python library specialized in Machine Learning. Because data read from the sensors is sequential, we’ve built our machine learning model based on Recurrent Neural Network and Long-short term memory structure because these types of neural network can process sequences of continuous data
We expect to achieve 90% of accuracy from the machine learning model.
Sampling Frequency: Several authors have suggested that the highest frequency components of the EMG signal are around 400-500 Hz [19], [20]. To follow the Nyquists rate, a smapling rate at least, 800-1000 Hz is needed to avoid aliasing and other signal distortion [21].
(https://core.ac.uk/download/pdf/96919379.pdf)

## Slide 9: ORIGINAL SCHEDULE

## Slide 10: UPDATED SCHEDULE

## Slide 11: TASK 1: Microcontroller Programming

## Slide 12: TASK 1: Test Plan and Result
Uniform Sampling: Check the number of samples per second
Bluetooth Connectivity: Expected range is 2-3 meters. Test to see how far the device can go before disconnecting

## Slide 13: TASK 2: Machine Learning Model

## Slide 14: TASK 2: Test Plan and Result
Test plan:
Test with separate datasets to check the accuracy.
Expected accuracy is 90% on the testing stage
Performance:
2,033,500 data points recorded
Training Accuracy: 88.97%
Testing Accuracy: 91.68%

## Slide 15: TASK 3: Smartphone Application

## Slide 16: TASK 3: Test Plan and Result
Data Collection and Visualization
Data Saver Function: Check to see if data is properly and efficiently stored in local memory
Test the model deployment on a subject
Performance:
Accuracy on the test subject: 70%

## Slide 17: PROJECT DEMONSTRATION
EMG Device | 05/03/2021

## Slide 18: CONSTRAINTS
EMG Device | 05/03/2021

## Slide 19: RISKS MANAGEMENT
High-risk Task:
Task 2.2 (ML Algorithm): High sensitivity to changes
  Mitigation Plan: improve the accuracy by adding more samples to the training dataset, generalize the input sequence, and tuning up the training parameters of the model.
 |  |  |  | Task 2.2
 |  | Task 2.4 | Task 2.3 | 
Consequence
Probability
1
2
3
4
5
1
2
3
4
5
EMG Device | 05/03/2021
Notes: training Parameters:
Number of Layer in the Neural Network
Number of Neurons in each layer
Learning rate
Number of epochs to train the model,

## Slide 20: RISKS MANAGEMENT
Medium-risk Tasks:
Task 2.3 (Training dataset): The credibility of the dataset is affected while collecting
  Mitigation Plan: Record samples under various circumstances so that the Machine Learning Model can learn about the differences
 |  |  |  | Task 2.2
 |  | Task 2.4 | Task 2.3 | 
Consequence
Probability
1
2
3
4
5
1
2
3
4
5
EMG Device | 05/03/2021
Notes: training Parameters:
Number of Layer in the Neural Network
Number of Neurons in each layer
Learning rate
Number of epochs to train the model,

## Slide 21: RISKS MANAGEMENT
Low-risk Tasks:
Task 2.4 ( ML Model Tune-up): Constrained by personal computer’s computability
  Mitigation Plan: Use cloud computing instead of personal computers.
 |  |  |  | Task 2.2
 |  | Task 2.4 | Task 2.3 | 
Consequence
Probability
1
2
3
4
5
1
2
3
4
5
EMG Device | 05/03/2021

## Slide 22: BUDGET
EMG Device | 05/03/2021

## Slide 23: EXPENDITURES
EMG Device | 05/03/2021
Parts | Cost
Microcontroller ESP32 (x2) | $23.80
MyoWare Muscle Sensor (x3) | $123.24
Non-Gelled Reusable Electrodes (x10) | $16.24
Gelled Electrodes (x100) | $19.14
Conductive Adhesive Gel (x1) | $5.72
Total | $188.14

## Slide 24: SUMMARY
Major Accomplishments:
Setbacks:
Lessons Learned:
EMG Device | 05/03/2021

## Slide 25: ACKNOWLEDGEMENT
EMG Device | 05/03/2021
Special thanks to our support:
Sponsor: Dr. Rose T. Faghih
Mentor: Md. Rafiul Amin
Project Manager: Dr. Steven Pei
University of Houston: Computational Medicine Lab

## Slide 26: Thank you!
Questions?
EMG Device | 05/03/2021

## Slide 27: SUGGESTIONS
3D unobtrusive wearable
Battery (18650 rechargeable battery 3.7V-2500 [mAh])
Larger dataset
Better sensors
Cloud database

## Slide 28: Facial Electrode Locations
Electrode locations for measuring facial EMG activity
EMG Device | 05/03/2021

## Slide 29: REFERENCES
https://www.therecoveryvillage.com/mental-health/ptsd/related/ptsd-statistics/
https://afsp.org/story/cdc-releases-new-mortality-data-nation-s-largest-suicide-prevention-organization
https://www.dbsalliance.org/education/depression/statistics/
https://www.scottdmiller.com/the-failure-rate-of-psychotherapy-what-it-is-and-what-we-can-do/
A. Van Boxtel, “Facial EMG as a Tool for Inferring Affective States,” vol. 2010, no. Figure 3, pp. 104–108, 2010.
EMG Device | 05/03/2021